QMS gap report - 20261026-2243

Model: qwen2.5:7b-instruct | Embeddings: sentence-transformers/all-MiniLM-L6-v2 | top_k=4, chunk=120 words, threshold=0.3

Draft for review by a qualified person. Not a compliance decision.

(a) Regulatory compliance strategy: WEAK

The extracts address the requirement partially, as they include risk management and supplier control, but do not explicitly mention a strategy for regulatory compliance or procedures for managing modifications.

Suggested fix: The company should develop a clear strategy for regulatory compliance and procedures for managing modifications to devices, in line with EU MDR requirements.

Cited: SOP-007_Design_and_Change_Control > 3. References; SOP-007_Design_and_Change_Control > 5. Procedure > 5.3 Risk management

Retrieved: SOP-007_Design_and_Change_Control > 3. References (0.395); SOP-004_Supplier_Control > 5. Procedure > 5.4 Incoming inspection (0.384); SOP-007_Design_and_Change_Control > 5. Procedure > 5.3 Risk management (0.379); SOP-004_Supplier_Control > 5. Procedure > 5.1 Selection (0.366)

(b) General safety and performance requirements: COVERED

The procedures mention risk management and field safety corrective actions, which align with the general safety and performance requirements.

Cited: SOP-007_Design_and_Change_Control > 5. Procedure > 5.3 Risk management; SOP-008_PMS_and_Vigilance > 5. Procedure > 5.5 Field safety corrective actions; SOP-007_Design_and_Change_Control > 3. References

Retrieved: SOP-007_Design_and_Change_Control > 5. Procedure > 5.3 Risk management (0.411); SOP-008_PMS_and_Vigilance > 5. Procedure > 5.5 Field safety corrective actions (0.406); SOP-007_Design_and_Change_Control > 3. References (0.377); SOP-005_Training_and_Competence > 1. Purpose (0.377)

(c) Management responsibility: COVERED

The extracts clearly define the responsibilities related to management review and the quality management system.

Cited: SOP-006_Management_Review > 4. Responsibilities

Retrieved: SOP-006_Management_Review > 4. Responsibilities (0.504); SOP-005_Training_and_Competence > 4. Responsibilities (0.458); SOP-002_CAPA > 4. Responsibilities (0.446); SOP-004_Supplier_Control > 4. Responsibilities (0.416)

(d) Resource and supplier management: COVERED

The extracts clearly outline the responsibilities and scope for supplier management, addressing the requirement.

Cited: SOP-004_Supplier_Control > 4. Responsibilities; SOP-004_Supplier_Control > 2. Scope; SOP-004_Supplier_Control > 1. Purpose

Retrieved: SOP-004_Supplier_Control > 4. Responsibilities (0.544); SOP-004_Supplier_Control > 2. Scope (0.501); SOP-004_Supplier_Control > 1. Purpose (0.466); SOP-004_Supplier_Control > (intro) (0.456)

(e) Risk management: COVERED

The procedures clearly outline the risk management process in accordance with ISO 14971:2007, and specify that the risk management file is updated at each design review.

Cited: SOP-007_Design_and_Change_Control > 5. Procedure > 5.3 Risk management

Retrieved: SOP-007_Design_and_Change_Control > 5. Procedure > 5.3 Risk management (0.564); SOP-007_Design_and_Change_Control > 3. References (0.447); SOP-002_CAPA > 5. Procedure > 5.3 Action plan (0.408); SOP-007_Design_and_Change_Control > 5. Procedure > 5.6 Design changes (0.392)

(f) Clinical evaluation: COVERED

The SOP-008 clearly outlines the preparation of a Post-Market Surveillance Report that includes updates to the clinical evaluation report, addressing the requirement.

Cited: SOP-008_PMS_and_Vigilance > 5. Procedure > 5.3 Post-market surveillance report

Retrieved: SOP-008_PMS_and_Vigilance > 5. Procedure > 5.3 Post-market surveillance report (0.493); SOP-002_CAPA > (intro) (0.39); SOP-005_Training_and_Competence > (intro) (0.375); SOP-008_PMS_and_Vigilance > (intro) (0.372)

(g) Product realisation: COVERED

The extracts cover product realisation by addressing design and development processes, including planning and change control.

Cited: SOP-007_Design_and_Change_Control > 2. Scope; SOP-007_Design_and_Change_Control > 5. Procedure > 5.1 Design planning

Retrieved: SOP-007_Design_and_Change_Control > 2. Scope (0.35); SOP-004_Supplier_Control > 2. Scope (0.273); SOP-007_Design_and_Change_Control > 5. Procedure > 5.1 Design planning (0.268); SOP-007_Design_and_Change_Control > (intro) (0.263)

(h) UDI: MISSING

The extracts do not mention UDI or Article 27 and 29 of the EU MDR.

Cited: SOP-007_Design_and_Change_Control > 5. Procedure > 5.4 Verification and validation; SOP-007_Design_and_Change_Control > 3. References; SOP-001_Document_Control > 3. References; SOP-002_CAPA > 5. Procedure > 5.4 Implementation

Retrieved: SOP-007_Design_and_Change_Control > 5. Procedure > 5.4 Verification and validation (0.413); SOP-007_Design_and_Change_Control > 3. References (0.391); SOP-001_Document_Control > 3. References (0.374); SOP-002_CAPA > 5. Procedure > 5.4 Implementation (0.345)

(i) Post-market surveillance: COVERED

The extracts clearly outline the post-market surveillance plan, its implementation, and maintenance, including data sources and reporting responsibilities.

Cited: SOP-008_PMS_and_Vigilance > 5. Procedure > 5.1 Post-market surveillance plan; SOP-008_PMS_and_Vigilance > 4. Responsibilities; SOP-008_PMS_and_Vigilance > 5. Procedure > 5.3 Post-market surveillance report

Retrieved: SOP-008_PMS_and_Vigilance > 5. Procedure > 5.1 Post-market surveillance plan (0.716); SOP-008_PMS_and_Vigilance > 1. Purpose (0.632); SOP-008_PMS_and_Vigilance > 4. Responsibilities (0.62); SOP-008_PMS_and_Vigilance > 5. Procedure > 5.3 Post-market surveillance report (0.592)

(j) Communication: WEAK

The extracts do not specifically address communication with competent authorities, notified bodies, other economic operators, customers, and other stakeholders.

Suggested fix: The company should establish a dedicated procedure for handling communication with competent authorities, notified bodies, other economic operators, customers, and other stakeholders.

Cited: SOP-002_CAPA > 4. Responsibilities; SOP-003_Complaint_Handling > 4. Responsibilities; SOP-004_Supplier_Control > 4. Responsibilities; SOP-008_PMS_and_Vigilance > 4. Responsibilities

Retrieved: SOP-002_CAPA > 4. Responsibilities (0.37); SOP-003_Complaint_Handling > 4. Responsibilities (0.35); SOP-004_Supplier_Control > 4. Responsibilities (0.35); SOP-008_PMS_and_Vigilance > 4. Responsibilities (0.345)

(k) Vigilance: COVERED

The procedure extracts clearly outline the process for notifying the HPRA of field safety corrective actions before they are undertaken, which addresses the requirement.

Cited: SOP-008_PMS_and_Vigilance > 5. Procedure > 5.5 Field safety corrective actions

Retrieved: SOP-008_PMS_and_Vigilance > 5. Procedure > 5.5 Field safety corrective actions (0.549); SOP-008_PMS_and_Vigilance > 1. Purpose (0.535); SOP-008_PMS_and_Vigilance > 4. Responsibilities (0.513); SOP-008_PMS_and_Vigilance > 6. Records (0.491)

(l) CAPA: COVERED

The extracts clearly outline the responsibilities, scope, and procedures for implementing and approving CAPA actions, addressing the requirement.

Cited: SOP-002_CAPA > 4. Responsibilities; SOP-002_CAPA > 2. Scope; SOP-002_CAPA > 5. Procedure > 5.4 Implementation; SOP-002_CAPA > 5. Procedure > 5.3 Action plan

Retrieved: SOP-002_CAPA > 4. Responsibilities (0.752); SOP-002_CAPA > 2. Scope (0.742); SOP-002_CAPA > 5. Procedure > 5.4 Implementation (0.709); SOP-002_CAPA > 5. Procedure > 5.3 Action plan (0.701)

(m) Monitoring, measurement and improvement: WEAK

The extracts address parts of the requirement but do not cover the full scope of monitoring, measurement, and product improvement processes.

Suggested fix: The company should establish a comprehensive process for continuous monitoring, measurement, and analysis of product performance and quality system outputs, with clear actions for improvement.

Cited: SOP-006_Management_Review > 5. Procedure > 5.4 Outputs; SOP-006_Management_Review > 5. Procedure > 5.3 Data analysis

Retrieved: SOP-006_Management_Review > 5. Procedure > 5.4 Outputs (0.48); SOP-004_Supplier_Control > 5. Procedure > 5.5 Monitoring (0.454); SOP-006_Management_Review > 5. Procedure > 5.3 Data analysis (0.401); SOP-008_PMS_and_Vigilance > 4. Responsibilities (0.386)